3.3.0

16.12.1996

10        * Address of the cryosystem                --- Main ---

31        * Type of the cryosystem

13        * Before terminator for send (0:not, 10:lf, 13:cr)

10        * End terminator for send

0         * Before terminator for enter

10        * End terminator for enter

0         * Protocol

0         * Available (Bit 0:Term for RS232, Bit 1:Protocol for IEEE)

9600      * Baud-rate for Com (600..19200)           --- RS232, LPT ---

1         * Stop-bits for Com (1..2)

8         * Number of bits for Com (7..8)

N          * Parity-Bit set for Com (Y/N)

N          * Even parity for Com (Y/N)

0         * Time-out (0:no, 1:auto, 2-65000:ms)

120       * RTS-Delay [ms] (<0:FIFO,+Mode*1000)

0         * Wait time after send and enter [ms]      --- IEEE, RS232, LPT ---

0         * Wait time mode (0:no, 1:send, 2:enter, 3:1&2, 4:not read)

0         * Reserved

332       * Internal version

-  *------------------ Temperatures and modes (line 22) ----------------------

10.0      * Minimum temperature [K]                  --- temperatures ---

330.0     * Maximum temperature [K]

0.01      * Minimum delta-temperature for ramp [K/s]

1.0       * Maximum delta-temperature for ramp [K/s]

0.5       * Offset stop temperature for ramp   [K]

0         * Reserved

1         * Number of temperature sensor             --- temp modes ---

0         * Numbers of sensors (0:no choice)

2         * Meas mode (2:normal, 3:1+2, 4:2+1)

0         * Mode for temp 2 (0:normal, 1:Pt, 2:extern, 3:Pt extern)

0         * Cryo-Pt/Pt file number (0:no)

0         * Work with temp. depending cmds

0         * Set temp at tempscan (0:std, 1:special, 2:1&prepare)

0         * Type of switching the cryo system        --- cryo modes ---

4         * Control avail

1         * Temp format (1:K, 2:K exp, 3:K/10, 4:C, 6:Ohm, 7:C/10)

0         * Answer number (0:last)

N          * Read answers at setting

0         * Ramp delta mode

0         * Ramp end mode (Bit 0:stop, 1:del, 2:set)

0.0       * Cable resistance [Ohm]

0         * Reserved

0         * Reserved

0         * ErrorChar                                --- errors ---

0         * ErrorMode (0:no, 1/2:first/last, 3/4:sig, 5:every)

1         * Error send loops

1         * Error enter loops

0         * Error bit at polling (0:no check, 1..8:error bit)

N          * Check protocol at reading

4         * Repeat at error (0:no, 1,2:no time out, 3-5:yes)

3         * Switch error loops

-  *------------------ Commands (line 54) -----------------------------------

0         * Load special overlay                     --- params ---

* Name of special overlay

Y          * Comments in cmds available

\          * Command character

0         * Reserved

S \T0 K                           * Cmd for set temp       --- base cmds ---

W0                                * Cmd for read sensor 1

* Cmd for read sensor 2

* Cmd for switch off     --- switch cmds ---

* Cmd for switch on

* Cmd for pump off

* Cmd for pump on

* Cmd for heater off

* Cmd for heater on

* Cmd for remote mode

* Cmd for local mode (ManualRamp)

* Cmd for UncontrolMode

* Cmd for ControlMode

* RampMode               --- ramp cmds ---

* Cmd for HoldMode

* Cmd for UnholdMode

* Cmd for end of ramp

* Reserved

* Ramp start, stop, delta temp

* 2. ramp cmds

* Change ramp

* Change ramp steps

* Error cmd, local mode  --- error cmds ---

* Error off, local mode

* Error cmd, remote mode

* Error off, remote mode

\IC\|\IR                          * Cmd for repeat, time out

* Cmd for repeat, no time out

* Cmd for get set point  --- other cmds ---

* Cmd for prepare tempscan

* Cmd for set at tempscan

-  *------------------ Extented params and commands (line 91) ---------------

0         * Version (0: no use)

-  *------------------ Temperature depending commands (line 93+X) -----------

0         * Numbers of following temp. depending cmds (0-6)

-  *------------------ Initialization commands (line 95+X) ------------------

0         * Numbers of the following init cmds (0-50)

*****************************************************************************

*****************************************************************************

*         Documentation for cryo system  (example type=31, DRC91)           *

*****************************************************************************

Line:

1) File version, dont change this line!

2) Date and comment

17) RTS-Mode (0:std, 2:half time at send/read)

38) Temp format (+100: delete blancs)

60) Command for setting of the temperature

61) Command for reading of the temperature

97+X) Init-Command at start of the program

******************************************************************************

*                     General explaination                                   *

******************************************************************************

**************************** Internal vriables *******************************

T0: Temperature set point [K]

T1: current temperature at standard sensor [K]

T2: current temperature at 2. sensor [K]

T3: Start temperature of ramp [K]

T4: Stop temperature of ramp [K]

T5: Delta value for ramp between start and stop temperature

T5: Delta value for ramp between current and stop temperature

**************************** Ports *******************************************

For the parameter of the IEEE-bus, the parallel port or the serial (Com-) port

is the same description valid as in the file HARDCONF\PERI_X.DAT.

The lines 11 until 17 are only valid for a serial port.

************************** Command syntax ************************************

The command character will be defined in line 58. Normally this character is

"\". A command starts with this character. Many commands in one line are

possible. The different commands must be separated by "\|".

In the following "cmd" means a command.

There are some groups of commands:

I) Transfer commands:

1) \S cmd  : Send a cmd

2) \E      : Enter an answer

3) \B cmd  : Send a cmd and enter the answer

4) \FX     : Define the protocol X (0:no secondary)

5) \sX cmd : Send a cmd with the secondary address X

6) \eX     : Enter an answer with the secondary address X

7) \bX     : Send a cmd and enter an answer with the secondary address X

II) IEEE commands:

1) \IC     : IFC, interface clear

2) \ID     : SDC, select device clear

3) \IG     : GTL, go to local

4) \IL     : LLO, local lock out

5) \IR     : REN, remote enable

III) Temperature commands:

1) \T0     : Value for set temperature

IV) Hardware commands:

1) \HHX    : Switch heater to mode X (0:off, 1:on)

2) \HPX    : Switch pump to mode X (0:off, 1:on, 8:on defined)

3) \HSX    : Switch full system to mode X (0:off, 1:on)

V) Special commands:

1) \CX     : CHR(X)

2) \VSX    : Set cryo sensor number to X, number starts at 1

3) \VCX    : Set cryo sensor number to X, number starts at 0

4) \WX     : Wait for x ms

VI) Commands for external program:

1) \LX     : Load external program with mode X

2) \LSX    : Load external program with mode X and send answer string

********************* Temperature depending commands *************************

In line 94+X you can define, whether you want to work with temperature

depending commands. M is the number of the command lines. Zero means no working

with temperature depending commands. In the other case you have 5 temperature

ranges. The input is:

Cmd line 1, range 1

......

Cmd line M, range 1

High temperature limit of range 1

Cmd line 1, range 2

......

Cmd line M, range 2

High temperature limit of range 2

.....

Cmd line 1, range M

......

Cmd line M, range M

Every cmd line start with the cmd line mode. Then follow the cmd(s) descripted

above. Following modes for the cmd lines exist:

0) No special mode, cmd will be done every time

1) For the temperature setting with steps bigger than 10 K and temp. decrease

2) For the temperature setting with steps bigger than 10 K and temp. increase

3) For the temperature setting with steps smaller than 10 K and temp. decrease

4) For the temperature setting with steps smaller than 10 K and temp. increase

5) For the temperature ramp, temperature decrease

6) For the temperature ramp, temperature increase

********************* Working with two sensors *******************************

Line 31 / Meas mode:

2) normal, temperature measurement with one sensor, control and measure with

sensor 1, temperature file data from sensor 1

3) temperature measurement with two sensors, control and main measure with

sensor 1, temperature file data from sensor 1, second data temperature

from sensor 2

4) temperature measurement with two sensors, control with sensor 1,

temperature file data from sensor 2, second data temperature from

sensor 1

For the use of this option you need normally a special hardware in your

cryo system

************************** New Biorad-controller *****************************

If you want an initialisation for the new Biorad-controller (cryosystem type

51 and 52) at start of the program (for example to set the PID-parameter),

so must be in the command line first the secondary address, then a space and

then the command. For example, the following command choose the PID-propor-

tional part: \s2 25

The integral part have the secondary address 12, the differential part the

address 13. The parameter must be an integer value between 0 and 100.

******************************************************************************

*                 Working with external EXE files                            *

******************************************************************************

It is possible to call external EXE files for reading temperature, setting

temperature, setting ramp and so on. It is always possible to combine the

normal work with this mode. For example you can read the temperature with

the standard program and set/start the ramp with an external program.

Data between the standard program and the external propram will be transfered

by the file TEMP.DAT. The subdirectory for this file is the buffer directory

BUFFER. If you work with a ramdisk this subdirectory will be on the ramdisk.

This file is an ASCII file. The format is:

I) standard program  -->  external program

1) Mode  (unsigned integer as ASCII)

2) Value (real as ASCII)

3) current TempStart, only modes 13 until 21

4) current TempEnd, only modes 13 until 21

5) current TempDelta, only modes 13 until 21

II) external program  -->  standard program

1) dummy flag

2) answer

The following modes exist:

0) Read temperature of sensor 1

1) Set temperature

2) Read temperature of sensor 2

3) Switch cryo system off

4) Switch cryo system on

5) Switch pump off

6) Switch pump on

7) Switch heater off

8) Switch heater on

9) Set cryo mode to Remote

10) Set cryo mode to Local

11) Set cryo mode to UnControl

12) Set cryo mode to Control

13) Set cryo mode to Ramp

14) Set cryo mode to Hold

15) Set cryo mode to UnHold

16) Set cryo mode to RampEnd

18) Set ramp start temperature

19) Set ramp start temperature 2

20) Set ramp end temperature

21) Set ramp delta temperature

22) Error cmd, local mode

23) Error off, local mode

24) Error cmd, remote mode

25) Error off, remote mode

26) Cmd for repeat, time out

27) Cmd for repeat, no time out

28) Get set point

31-36) Temperature depending commands 1-6, range 1

37-42) Temperature depending commands 1-6, range 2

43-48) Temperature depending commands 1-6, range 3

49-54) Temperature depending commands 1-6, range 4

55-60) Temperature depending commands 1-6, range 5

******************************************************************************

*                            Cryo systems                                    *

******************************************************************************

Specifig types of cryo systems with CRYO_*.DAT-files (file/type-number):

- / 0: no cryo system, simulation

- / 1: no cryo system, only room temperature

- / 2: no cryo system, room and N2 temperature

11 /11: external Keithley resistance meter and Pt-table

21/ 21: external EXE-file

22 /22: old Cryophysics-He (as DRC91/old but with waiting time)

23 /23: temperature controller T-3000

24 /24: MMR K-20

25 /25: Eurotherm-RS232, Celcius

31 /31: Cryophysics-He DRC91/old, old Biorad-He

32 /32: Eurotherm-RS232, Kelvin

33 /33: Oxford ITC503

34 /34: temperature controller 9650

35 /35: Cryophysics-He DRC93/91

36 /36: Oxford ITC502,ITC4

41 /41: Biorad-N2 old

51 /51: Biorad-N2 new

52 /52: Biorad-He new

53 /53: AME

54 /54: Chinon

55 /55: Linkam controller old

56 /56: Linkam controller new

Global types of cryo systems:

X: no cryo system, without measurement of temperature

1X: without set of temperature

2X: measure and set of temperature

3X: measure and set of temperature

4X: temperature-ramp manuell possible

5X: temperature-ramp computer-controlled possible

Switching types of the cryo system:

1: heater possible   2: pump possible

3: 1 & 2             4: all

5: all and heater    6: all and pump

7: 3 & 4